## Slide 1: Directionality Characteristics of the Tympan Open-Source Hearing Aid and Earpieces
Benjamin P. Masters
Joshua M. Alexander

## Slide 2: Methods
When possible, followed procedures outlined in ANSI/ASA S3.35-2021
  “Method of measurement of performance characteristics of hearing aids under simulated real-ear working conditions”
Pink Noise Stimulus
  ≈85 dB SPL at hearing aid mic location
  5-second duration
  Analyzed in 1/3-octave bands

## Slide 3: Set up
KEMAR centered in a double-walled sound booth (2.13 m x 2.44 m)
  Rotated in 15° increments
  Left earpiece was tested
Hafler M5 Reference Monitor sound field loudspeaker
  Flat frequency response
3
Test ear

## Slide 4: Earpieces
9.5-mm mic spacing (d)
  c ≈ 343 m/s
  Travel time (d/c): 27.7 μs
Sampling rate set to 36.1 kHz → 1-sample delay = d/c = 27.7 μs
  Cardioid directivity
9.5 mm
27.7 μs
4

## Slide 5: Microphone Matching
Critical for optimal cancellation
  Rear mic output was ≈ 5 dB lower than the front mic
    5 dB of gain was added to match the two mics

## Slide 6: Omni-Directional Responses
It is well-known that putting mics on the head will shift the response to the side due to head shadow
AIDI = -1.12
AIDI = -1.00
Front Mic (re 0°)
Rear Mic (re 180°)

## Slide 7: Rear Mic Delay
AIDI = +2.31
(+3.43)
AIDI was lowest with -6 dB and +6 dB rear mic gain
AIDI = +0.50
(+1.62)
AIDI = -1.53
(-0.41)
0° reference
-3 dB Rear Gain
0 dB Rear Gain
+3 dB Rear Gain
*Gain values after mic matching
AIDI = Articulation Index weighted Directivity Index values in () = Improvement re: front mic only

## Slide 8: Front Mic Delay
180° reference
-3 dB Front Gain
0 dB Front Gain
+3 dB Front Gain
*Gain values after mic matching
AIDI was lowest with -6 dB and +6 dB front mic gain
AIDI = Articulation Index weighted Directivity Index values in () = Improvement re: rear mic only
AIDI = +0.95
(+1.95)
AIDI = +2.98
(+3.98)
AIDI = -0.11
(+1.11)

## Slide 9: AIDI Improvement
Front mic with rear mic delay and gain = -3 compared to front mic only (0° reference)
Rear mic with front mic delay and gain = 0 compared to rear mic only (180° reference)

## Slide 10: Future Directions: Variable Directionality
(Kates 2008)
DELAY
d/c sec
DELAY
d/c sec
∑
REAR
MIC
FRONT
MIC
d
∑
-
-
+
+
∑
+
-
GAIN
b
Step 1: create opposing cardioid patterns
  Front facing
  Rear facing: inverted phase and variable gain, b
Front Cardioid
Rear Cardioid
d = mic separation
c = speed of sound

## Slide 11: Future Directions: Variable Directionality
b = 1
b = 0.5
b = 0.25
The front and rear responses add with opposite sign, thereby creating nulls where they intersect
Step 2: vary relative gain b on rear cardioid to create different polar patterns with a continuous range of nulls
Bidirectional
Hypercardioid
Supercardioid

## Slide 12: Conclusions
Optimized directionality requires precise timing (sampling rate) and microphone matching
Directional patterns and relative AIDI improvements (≈3-4 dB) are consistent with published research
Optimal directivity
  Front cardioid obtained with -3 dB gain on the rear mic
  Rear cardioid obtained with matched mic gain
Future directions: use optimized front and rear cardioids with opposite phase to vary null with fixed and adaptive directionality